Mailing Lists: I have completed updating the mailing list. I thank everyone who returned the forms to me.

Costs: At PTCOG XIX, the Steering Committee decided that part of the registration fee for PTCOG meetings would be used to help produce both Particles and the abstracts of the PTCOG meetings. Only part of the costs are covered in this way, so more financial help is needed from the community. We are always happy to receive financial gifts

HCL is always happy to receive financial gifts; all such gifts are deductible as charitable contributions for federal income tax purposes. The appropriate method is to send a check made out to the ÒHarvard Cyclotron LaboratoryÓ. We thank Dr. Steven Goetsch for his kind contribution.

Facility and Patient Statistics: I continue to collect information about all operating or proposed facilities. Please send me your information. My most recent published summary of the world wide patient statistics with detailed patient data through 1994 can be found in the following reference. ÒProton therapy in 1996.Ó J. M. Sisterson, CP392, Application of Accelerators in Research and Industry, eds. J.L. Duggan and I.L. Morgan, AIP Press, New York (1997), p1261-4.

Particles on the Internet: We have set up a home page for the Harvard Cyclotron Laboratory on the Internet with links to recent issues of Particles.

The URL for the Harvard Cyclotron Laboratory is:-

http://neurosurgery.mgh.harvard.edu/hcl/        or        http://brain.mgh.harvard.edu:100/hcl

Other proton therapy links: (I did try all these URLs but I am sure this list is not complete, so PLEASE send me your URL to include in the next issue).

Northeast Proton Therapy Center: http://www.mgh.harvard.edu/nptc/nptc.htm

LLUMC, California: http://www.llu.edu/proton

U of California, Davis: http://crocker.ucdavis.edu/cnl/research/eyet.htm

Indiana University: http://nike.iucf.indiana.edu/ptherapy/

TRIUMF, Canada protons: http://www.triumf.ca/welcome/proton_thrpy.html

TRIUMF, Canada pions: http://www.triumf.ca/welcome/pion_trtmt.html

NAC, South Africa: http://www.nac.ac.za/~medrad/index.html

KVI, The Netherlands: http://www.kvi.nl/disk$1/protonlib/www/homepage.html

PSI, Switzerland: http://www1.psi.ch/www_asm_hn/asm_home_page.html

Proton Oncological Therapy, Project of the ISS, Italy: http://top.iss.infn.it

TERA foundation, Italy: http://www.tera.novara.it

Tsukuba, Japan: http://www-medical.kek.jp/index.html

Tsukuba, Japan - new facility plans: http://www-medical.kek.jp/devnewfac.html

HIMAC, Chiba, Japan: http://www.nirs.go.jp/ENG/particl.htm

National Association for Proton Therapy: http://www.proton-therapy.org

Prolit - database of particle radiation therapy; http://proton.llu.edu

GSI homepage: http://www.gsi.de

ARTICLES FOR PARTICLES 23

The deadline for news for Particles 23, the January 1999 issue, is November 30 1999. I will send reminders by fax or e-mail.

Address all correspondence for the newsletter to:

Janet Sisterson Ph.D.	Telephone: (617) 495-2885

Harvard Cyclotron Laboratory	Fax: (617) 495Ð8054

44 Oxford Street	E-mail: 	sisterson@radonc.mgh.harvard.edu

Cambridge MA 02138

Articles for the newsletter can be short but should NOT exceed two pages in length. The best way to send an article is by computer. If you mail or fax an article, remember that I scan them into the computer so I need a good clean copy of any figures.

PTCOG and FUTURE PTCOG MEETINGS

| Chair: Michael Goitein | Secretary: Janet Sisterson |
| --- | --- |

Department of Radiation Oncology

Massachusetts General Hospital

Boston MA 02114

Harvard Cyclotron Laboratory

44 Oxford Street

Cambridge MA 02138

The PTCOG e-mail address is PTCOG@radonc.mgh.harvard.edu

It is with regret that we accept Dr. Joe CastroÕs resignation from the Steering Committee. We are pleased that Dr. Herman Suit has agreed to take his place on this committee.

Steering Committee Members

| USA | Europe | Russia | Japan | South Africa |
| --- | --- | --- | --- | --- |
| W. Chu | U. Amaldi | V. Khoroshkov | K. Kawachi | D. Jones |
| M. Goitein | H. Blattmann |
| H. Tsujii |
|  | D. Miller | J.-L. Habrand |
|  | J. Sisterson | G. Munkel |
|  | James Slater | E. Pedroni |
|  | A. Smith | A. Wambersie |
|  | H. D. Suit |
| L. Verhey |

The times and locations of the next PTCOG meetings are as follows:-

PTCOG XXVIII

Loma Linda, CA USA

April 15 - 17 1998

PTCOG XXIX

Heidelberg, Germany

September 14- 16 1998

PTCOG XXX

NAC, Cape Town, South Africa

April 12 - 15 1999

PTCOG XXXI

Bloomington, IN, USA

October 11 - 13 1999

The proceedings for the PTCOG XXVII meeting held in Chiba, Japan in November 1997 should be available soon.

SECOND ANNOUNCEMENT:

PTCOG XXIX,

Heidelberg / Germany,

September 14 - 16, 1998.

Outline of the Preliminary Program

Sunday, 13th at the Renaissance Hotel, Heidelberg

18:00 - 20:00	Welcome Reception

Monday, 14th

9:00 -17:45 	Sessions at the DKFZ, Heidelberg

19:00 - 22:00	Social Event

Tuesday, 15th.

7:30 - 8:20	Steering Committee Meeting breakfast

8:30 	 Transfer to GSI.

10:00 - 17:30 	Sessions at GSI, Darmstadt

Barbecue at GSI, then transfer to hotels.

Wednesday, 16th

9:00 - 12:30 	Sessions at the University Hospital, Heidelberg.

14:00-17:00	 meeting of the ACT

DKFZ, The University Hospital and the Renaissance Hotel are all in Heidelberg. GSI, Darmstadt is about one hour away.

Registration:  The registration form included in this mailing should be sent to:

dkfz

PD Dr. Dr. Jürgen Debus

Im Neuenheimer Feld 280

69120 Heidelberg

Germany

Accommodation: Participants are expected to make their own reservations until August 10th 1998 at (please note space is limited):

RENAISSANCE HEIDELBERG HOTEL

Vangerwostr. 16

69115 Heidelberg

Tel: + 49 6221 908-0

Fax + 49 6221 908 508

(Please refer to the above meeting in order to obtain special rate of about 200 DM)

Information in other accommodations can be obtained by Heidelberg tourist service: 	Telephone	+ 49 6221 19433

+ 49 6221 142223 / 24

Fax 	+49 6221 167318

Registration fee: The registration fee for the meeting has been set at DM 350.--. DM 300 .-- if paid before August 15th 1998. (DM 300.-- for those registered before June 30th 1998)

Please transfer the fee to one of the following accounts with the code H 206

Postscheckamt Karlsruhe ( 660 100 75 ) 452 75 - 752
Landeszentralbank Heidelberg ( 672 000 00 ) 672 / 01900
Deutsche Bank Heidelberg ( 672 700 03 ) 01 / 57008
Dresdner Bank Heidelberg ( 672 800 51 ) 4 688 491

Abstracts: Contributors are invited and strongly encouraged to submit an abstract of their

presentation that will be published in the January 1999 issue of Particles. The abstract should be about one half page in length, include authors and affiliations. Abstracts will be collected at the meeting or may be sent to Janet Sisterson by one of the means listed earlier in this newsletter. THE BEST METHOD is by e-mail to Sisterson@radonc.mgh.harvard.edu.

COMBINED MEETING    -    PTCOG XXX / EHTG / ECHED

12-15 April 1999

PTCOG			Proton Therapy Co-Operative Group

EHTG			European Hadron Therapy Group

ECHED			European Clinical Heavy Particle Dosimetry Group

Date:				12-15 APRIL 1999

Venue:			ARTHURÕS SEAT HOTEL

SEA POINT, CAPE TOWN

SOUTH AFRICA

Organisers:		Medical Radiation Group			Contact:

National Accelerator Centre		Dr Dan Jones

P O Box 72					Tel:  +27-21-843-3820

Faure					Fax: +27-21-843-3382

7131 SOUTH AFRICA			e-mail: jones@nac.ac.za

Latest Information:  http://www.nac.ac.za/~medrad/events.htm

The meeting will cover all aspects of neutron capture, fast neutron, proton and heavy ion therapy including

Clinical results

New treatment protocols

Accelerators

Beam delivery

New facilities

Dosimetry

Quality Assurance

Radiobiology

Treatment Planning

Cape Town is regarded as one of the worldÕs most beautiful cities and April is a very pleasant time of the year - the average maximum temperature is 23¡C/73¡F and the average minimum temperature is 12¡C/54¡F.  The ArthurÕs Seat Hotel is conveniently located in the suburb of Sea Point and is one block from the seafront.  Downtown is easily accessible as the hotel is on the main bus and taxi  routes and there is also a shuttle bus to the Waterfront shopping, hospitality and entertainment complex.  Cape Town is easily reached from all major  cities, either on South African Airways or on other international carriers.  Flights terminate in Cape Town or Johannesburg, which  is a 2-hour  flight  from  Cape Town.

For  information on siteseeing and tours contact the nearest office of the South African Tourism Board (SATOUR) or:

Carol Swart				Tel: +27-21-797-1861

Talk Travel				Fax: +27-21-797-7810

P O Box 6878

Roggebaai

8012  SOUTH AFRICA

If you  wish to receive further information please complete the form included in this mailing, and return it to the organisers.

BEAM SCANNING WORKSHOP

Rancho Mirage CA, 13-14 April 1998

Systems for scanning charged particle beams are complex, costly and demanding in terms of equipment, manpower and time.  Several groups are presently interested in developing such systems while some expertise in this area is already available from other groups.  The Beam Scanning Workshop brought together engineers, technical experts, medical physicists and other experienced persons from both sets of groups for discussions of all aspects of beam scanning.

This Workshop was arranged as a result of favourable responses received to the proposal in PARTICLES No. 20 (July 1997) regarding the formation of a beam scanning collaborative group and was attended by 26 participants.  The proceedings were divided into several sessions covering different topics. Each topic was introduced by different speakers after which discussions took place.  Apart from the main agenda topics, Dr Y Futami of Chiba give a brief presentation of the C-11 radioactive beam spot scanning system under development at NIRS.

AGENDA

WELCOME	[D Jones, NAC]

SPECIFICATIONS FOR BEAM SCANNING	[M Goitein, NPTC]

PATIENT POSITIONING	[N Schreuder, NAC]

BEAM DELIVERY	[J Flanz, NPTC]

RADIOBIOLOGY	[G Kraft, GSI]

TREATMENT PLANNING	[E Pedroni, PSI]

CONTROL AND SAFETY	[C Bloch, IUCF]

DOSE VERIFICATION	[M Schippers, KVI]

ESTABLISHMENT OF COLLABORATIVE GROUPS	[D Jones, NAC]

INTRODUCTION

The advantages of beam scanning systems in comparison with passive beam spreading systems are well known;  perhaps the most important are:

Intensity modulation (and inverse planning) is possible.

There is negligible reduction in the range of the beam.

Integral dose is reduced as dose conformation to the proximal edge of the lesion is possible.

In principle no field-specific modifying devices are required.

Scanning systems are completely flexible.

The main disadvantages include:

Scanning systems are more complicated and therefore potentially less reliable and more dangerous.

The development of such systems is more demanding in terms of cost, time and manpower.

More stable beams are required.

Dose and beam position monitoring are more difficult.

The problems associated with patient and organ movement are more severe.

There are several techniques which can be used for scanning.  For lateral beam spreading, circular scanning (wobbling) or linear scanning can be done.  In the latter case the beam can be scanned continuously or in a discrete fashion (spot scanning).  Another possibility is to undertake the fastest scan in one-dimension (strip scanning) and move the patient or the scanning magnet in the other dimension.  Depth variation is achieved by interposing degraders in the beam (cyclotrons) or by changing  the beam energy (synchrotrons).

The aim of beam scanning is to deliver a predetermined dose at any point in the body.  Special care must be taken as scanning (because of high instantaneous dose rates) is potentially more dangerous than passive beam modification systems.  The beam position and the dose delivered at each point must be accurately determined.

SUMMARY OF DISCUSSIONS

SOURCE-AXIS DISTANCE (SAD)

Cartesian scanning (ie. with infinite SADs  and therefore parallel beams) is most desirable as producing box-shaped fields and the patching together of adjacent fields (to treat larger areas) are in principle simpler.

With polar scanning (ie. with short SADs and therefore divergent beams) these procedures are more difficult but can be done with proper treatment planning.

With short SADs there is also an increase in the entrance dose (inverse square law), but the impact of this effect is reduced when multiple fields are used.

FIELD SIZE

Up to 40 cm in 1 dimension.

The length of the other dimension depends on the scanning technique used.  At least 40 cm is desirable.

The smallest field is determined by the elemental beam size.

DEPTH OF PENETRATION

A range of from near the skin surface (say, 1 or 2 cm.) to at least 32 cm.

RANGE MODULATION

For scanned beams, this concept doesnÕt apply.  However, the range of depths over which Bragg Peaks are applied is unlikely to be more than 20 cm proximal to the deepest depth required.  (This probably has little implication for design, given the previous specification.)

SPATIAL RESOLUTION (s) AND PENCIL BEAM SIZE

This specification cannot exceed the physical limit set by multiple scattering within the patient - which is depth dependant.  Typically, s near the end of range is of the order of 3% of the range (e.g. 3 mm at 10 cm depth).

s approx = 3 mm. at and near edge of field.

Within the field, a broader beam could be contemplated, say s = 6 mm or even more.

The sharper resolution at the field edge can be achieved: with adjustable pencil beam profiles; with a collimator; by phase space trimming; or, to some extent, by a variable intensity near the field edge.

Changing the spot size between layers to preserve spacing should be considered, but this effect can be taken into account in the treatment planning.

DOSE DYNAMIC RANGE (PER FIELD)

Up to 8:1 within a depth ÒlayerÓ.

Up to 100:1 between layers.

DOSE FIDELITY

The delivered dose should everywhere be equal to the intended dose to within ± 2% - or a point receiving the intended dose should be found within a distance of ± 1 mm of any point of interest.

BEAM SPACING

Less than one third of the FWHM (full- width at half- maximum) of the pencil beam size.

NO. OF FIELDS PER FRACTION DELIVERED

From 7 to 30 (to simulate arc therapy) from any arbitrary direction.

TIME TO DELIVER ONE FIELD / FRACTION

The time to deliver a field should not exceed from 20 to 60 secs - depending on the field size and number of fields per fraction.

The more important specification is the time to deliver one fraction.  Ideally, after patient set-up and localization, this should take no more than from 5 to 10 minutes.

POSITIONING AND IMMOBILIZATION

Patient positioning and immobilization are in principle no different for scanned beams than for conventional proton radiotherapy or radiosurgery, except that organ or patient motion pose more severe problems.

Patient set-up is also more difficult as no field-defining lights can be used for field shape  and position confirmation. In practice x-ray verification of position is almost always done.

Typical set-up accuracies which are needed are <0.2 mm (extremities), 1-2 mm (head and neck) and 5-10 mm (abdomen).

ORGAN AND PATIENT MOTION

If the lesion being irradiated (or the patient) moves during treatment, hot and/or cold spots within the tumour and/or outside the tumour can be produced.  These effects can be minimized by undertaking multiple scans, reducing the distance between spots or by using larger spots.

Gating the beam on and off at specific patient or organ positions can be done.Up to now this has usually been related to the breathing cycle.  This requires real-time monitoring of the organ or patient position by means of motion detectors and transducers.

Other motion detector possibilities include the use of internal or external fiducial markers or anatomical landmarks in conjunction with x-radiography, ultrasound, neutron radiography or externally applied magnetic fields.

Patients can also be required to hold their breath or the areas of their bodies being irradiated can be compressed.

Because patient and organ motion usually takes place in a regular and predictable fashion the extent of movement can be anticipated and allowed for with reasonable accuracy.

The capability of repositioning from outside the treatment room is desirable to allow adjustment for small changes in patient or target position.

SCANNING MAGNETS

A long throw (distance from scanning magnets to patient)  is desirable as this means that smaller magnets are required and faster scans are possible (for a given change in the magnetic field the spot will move a longer distance in the same time than with a short throw).

Long throws are possible with fixed beams but the throw is limited in a gantry.

It is desirable to sweep the beam by 20 cm in 30-100 msecs.

SHAPE OF BEAM SPOT

Planning is easier if the spot shape is always the same and for practical reasons it is assumed to be invariant.

A symmetrical shape is ideal.

Asymmetrical spots do not pose problems if the spot spacing is close enough.

SAFETY

It goes (almost) without saying that the safety of a scanning system (as with any part of a treatment apparatus) is of paramount concern.

For safety reasons, asymmetrical scanning is probably ideal ie. the scanning magnet axes are offset from the treatment central axis.

Verification of the treatment pattern should take place as close to the patient as possible.

Redundant checks of the beam position must be done, preferably  with different computers. eg. the position can be checked by feedback of magnet current and of magnetic field strength.

Great care must be paid to restart procedures should interruption occur during irradiation in order to reproduce the required dose distribution.

BIOLOGICAL EFFECTS

The RBE depends on particle type, beam energy (depth), type of irradiated tissue and the integral dose (the latter is a result of the shape of the photon [reference] survival curves).

At the dose rates used in beam scanning there are no dose-rate effects.

Long treatment times are undesirable because they may approach the time scale of repair mechanisms.

With C-12 beams at GSI the RBE is changed for each irradiated elemental volume.

Although the RBE for proton beams does increase as a function of increasing depth (decreasing energy) this is not usually taken into account with any form of beam delivery.

A global RBE (usually 1.10) is assumed,  but recent measurements at LLUMC, NAC and PSI done by the same group, using the same biological system in similar beams gave RBEs of 1.05, 1.15 and 1.25 respectively.

TREATMENT PLANNING

To date dedicated systems have been required.  It has not been practical to modify commercial photon planning systems.

Intensity modulated beam delivery has been implemented at PSI only.  Intensity modulated plans have been developed, and intensity modulated treatments are likely to be delivered in the near future.

A standard elemental beam shape is assumed for planning purposes.  How often this shape should be checked is not clear (between fields, every day, ?)

A basic problem is that all dose distribution measurements are made in homogenous phantoms which do not reflect the situation in a patient.

Planning optimization is usually done at present by iterative techniques.  At PSI, the treatment planning system produces a so-called Òsteering fileÓ which specifies the scanning parameters for a particular field.  For each spot the magnet currents, beam energy/degrader thickness, couch position (if applicable), RBE (if applicable) etc. are given).

In principle the RBE should be taken into account throughout the treated volume.

DOSE CALIBRATION AND VERIFICATION

The absolute dose must be determined at specific reference points in a phantom in order to calibrate the dose monitors.

The 3-D dose distribution must be verified.

Different detector systems can be used for simulations and for verification during treatment.

For simulation the following detectors can be used:

1-D:	diodes, ionization chambers, TLDs, diamond detectors

2-D:	Òmagic cubeÓ, radiographic film, scintillation screens

3-D:	GEL (dose accuracy:  5%, spatial resolution: 1.5 mm

PET (2%, 4 mm)

CCD cameras with scintillation screens (0.5% , 0.5 mm)

During treatment multiwire ionization chambers, radiographic film (not real-time) and scintillation screens can be used.

REQUIREMENTS FOR QUALITY ASSURANCE

In order to undertake quality assurance procedures knowledge of the planned dose distributions are required.

The dose distributions must be checked with independent systems under exactly the same conditions as pertain to the treatment.  It should be noted that for some detection systems this is not possible (eg.  radiochromic film requires higher doses, PET requires beam interruptions for data acquisition).

New dosimetry protocols will have to be developed for scanned beams.

WORKING GROUPS

At the end of the Workshop four specific topics were identified for further attention.  It was proposed that Working Groups should be established to examine these topics in detail and report back at forthcoming PTCOG meetings.

The subjects to be addressed by the Working Groups, their Chairmen and e-mail addresses are given below:

|  | Working Group |
| --- | --- |

Chairman

|  | e-mail address |
| --- | --- |
| 1 |

SPECIFICATIONS FOR IDEAL SCANNING SYSTEM

|  | Michael Goitein |
| --- | --- |

Goitein@hadron.mgh.harvard.edu

Michael.Goitein@psi.ch

2

ONLINE CONTROL AND MONITORING OF SCANNED BEAMS

|  | Jay Flanz | Flanz@hadron.mgh.harvard.edu |
| --- | --- | --- |
| 3 |

EFFECTS OF ORGAN MOTION

|  | Dan Jones | Jones@nac.ac.za |
| --- | --- | --- |
| 4 |

DOSIMETRY OF SCANNED BEAMS

|  | Marco Schippers | Schippers@kvi.nl |
| --- | --- | --- |

It is anticipated that additional informal discussions among interested people will take place at PTCOG XXIX (Heidelberg), while a second workshop will be held next at PTCOG XXX (Cape Town) which will give us more time to prepare the topics in more detail than can be done in the short period since the last workshop in Palm Springs.

Those interested in partaking in the activities of any of the above groups should contact the respective Chairmen as soon as possible.

Dan Jones, Marco Schippers and Michael Goitein,
13 July, 1998

Clinical Protocol Working Group

At the 28th PTCOG meeting in Rancho Mirage, the Clinical Protocol Working Group had itÕs first session.  The main objectives of this working group are:

to provide participating members and all other interested parties with a continued update and status report of ongoing trials, including accrual rate, estimated time of protocol duration, and update of any foreseen or unforeseen difficulties.

to discuss objectives and treatment rationales of potential future trials.

to encourage international cooperation, specifically participation of institutions and facilities who agree with the principle treatment concept, and share similar technological equipment, etc.

This first session focused primarily on the status report of all currently existing proton protocols. Several potential future trial designs in the area of prostate cancer, CNS, sarcoma, and pediatric oncology.  The reaction of participants was positive and encouraging.  All relevant facts were summarized to the full PTCOG audience, on the last meeting of the day.

We plan to continue the working group at the upcoming meeting in Heidelberg, as well as future meetings, similar to other multi institutional and international cooperative groups.

Emphasis of the upcoming Fall meeting in Heidelberg will be to discuss participation, and future trials initiated by non U.S. institutions.

In keeping with PTCOG tradition, the clinical protocol working group is open to everyone interested.  We plan again, to present a summary to the full PTCOG audience.

Many of us are unfamiliar with technical equipment, beam availability, patient availability, etc., of the various institutions.  These and other variables are all factors influencing the choice of disease, site, and design of new trials, as well as opportunities to participate.  Please consider a brief 5 - 8 min. presentation and outline of these parameters of your facility at the session in Heidelberg. Please keep in mind that this is not the forum to present any clinical data, which should always be presented at the regular clinical sessions.

Eugen B. Hug, M.D., Associate Professor,Radiation Medicine, Loma Linda University Medical Center, Loma Linda California. Phone: (909) 824-4280, Fax: (909) 824-4083, E-mail: Ehug@dominion.llumc.edu.

PROLIT: Culling The Heavy-Particle Radiation Therapy Literature

Prolit, a database of Medline abstracts related to particle radiation therapy, is now available on the Loma Linda Proton Treatment Center Web pages (http://proton.llu.edu). The database provides access to over 5000 particle therapy abstracts, enabling physicians, patients, and researchers to begin their search for information in one convenient site on the Web.  A full description of Prolit is available on the Web site. The database will be updated regularly.

The first Prolit database was developed at LLUMC in the late 1980s; it was distributed to PTCOG members and others interested in particle radiation therapy via hard copy and diskettes. Prolit was discontinued in 1993, when it became apparent that the data-collection process was too labor-intensive and the mode of distribution limited. The current incarnation of Prolit can be updated quickly, and should reach a much larger audience via the Internet.

The Prolit development team anticipates refining the search engine and PubMed search strategy based on user feedback. Please direct any comments to Robert Kirby at rkirby@dominion.llumc.edu. Robert Kirby, Dept. of Radiation Medicine, Loma Linda University Medical Center, 11234 Anderson Street, Loma Linda, CA 92354.

PTCOG Information/News/Reports:

The following reports and articles were received by July 1998.

The National Association for Proton Therapy:

PROTON-THERAPY WEB SITE: now available on the internet

The National Association for Proton Therapy (NAPT) announced on March 18, 1998 the launching of its new web site --http://www.proton-therapy.org --to meet increasing demand for information on proton therapy for cancer treatment.

The web site focuses on the most frequently asked questions about the benefits of proton beam technology in the treatment of numerous cancer sites in the body -- including results of a recent study on men with prostate cancer. The web site includes sections on how protons work, proton facts, where to find proton therapy treatment, treatment center floor plan, testimonials from patients, and links to other proton therapy sites.

The NAPT is a non-profit organization providing education and awareness for the public, professional and governmental communities. Founded in 1990, it promotes the therapeutic benefits of proton therapy for cancer treatment in the U.S. and broad.

Internet users are encouraged to visit the site and ask questions of their own by contacting lenarzt@proton-therapy.org. Contact: Len Arzt, National Association for Proton Therapy, 7910 Woodmont Ave. #1303, Bethesda, MD 20814; 301-913-9360.

Midwest Proton Radiation Institute, Bloomington, In, USA update:

Dr. Alejandro Mazal (CPO, Orsay) visited Indiana University for two months this spring to help us develop specifications for the new treatment rooms. While Ale was here, he also undertook to educate us on many aspects of proton therapy through informal talks on Stereotactic Irradiations, Dosimetry, Radioprotection, Treatment Planning, etc.

We have completed development and testing of the eye treatment facility and are now ready to treat patients, pending final approval from IU's Machine Produced Radiation Safety Committee. The testing process took somewhat longer than anticipated. The first AMD patients are scheduled to be treated in mid-July. These patients are participating in a clinical trial on choroidal neovascular membrane in age-related macular degeneration.

Design work for the new facility has been centered on energy selection and beam transport from the Cyclotron to the treatment rooms. An architectural firm is working on a conceptual design for the building addition required for patient care, treatment planning and preparation, and staff support. Moira Wedekind, Indiana University Cyclotron Facility, 2401 Milo B Sampson Ln, Bloomington, IN 47408.

News from Moscow, Russia:

The feasibility study of The Project of PTF at Moscow Oncologic Hospital #62 is completed by ITEP in collaboration with five physical and medical institutes in April, 1998. At the moment the facility cost estimated at  $M23.3 and is being covered by Moscow Local Government. Hospital based PTF is being built around H- synchrotron with the energy of 230 Mev and will be equipped with four treatment units - 2 gantries and 2 units for horizontal fixed beam.

Patient treatment has been carried out at ITEP PTF in 1997. 3039 patients were irradiated by January 1, 1998. V.S. Khoroshkov, Inst. for Theor. & Exper. Physics, B. Cheremushkinskaya 25, Moscow 117259, Russia.

Heavy-Ion Therapy at GSI, Darmstadt, Germany:

After the successful treatment of the first two patients in December last year, the accelerator and the experimental area have been shut down because of service and upgrating work at the accelerator SIS and at the storage ring ESR and the installation of a new physics area that was planned long before .

The follow up of the two treated patients is very satisfying. In the CT-images a reduced uptake of contrast drug is observed and the T2 weighted MRI does not show any abnormality outside the target area. Therefore, it can be concluded that the treated volume coincides precisely with the target volume. This coincidence was also measured online during the treatment using PET techniques.

Because of the shut down period, the changes in the beam transport system and the installation of the heavy new physics cave, causing sedimentation problems to the building, a longer period of test and realignment is expected before patient treatment will be possible in order to obtain the same quality of performance that was reached last time.

At present the results of these first treatments are implemented in the control and safetxy system as well s in the treatment planning procedures. Next patient treatments are in preparation for late of August and early September. Then a maximum number of ten patients will be treated with twenty fractions in twenty consecutive days including weekends.

On September 14 -16 the next PTCOG meeting will be held at Heidelberg combined with the inauguration ceremonies for the GSI therapy.

More information is available on the web on the GSI homepage: http://www.gsi.de. Gerhard Kraft, GSI, Planckstr. 1, Darmstadt D 64291, Germany.

News from NLHIAL(National Laboratory of Heavy Ion Accelerator, Lanzhou), Lanzhou, P.R.China: HIRFL-CSR PROJECT IN LANZHOU

The Heavy Ion Research Facility in Lanzhou (HIRFL) has opened a new field of heavy ion physics and its application at intermediate energy domain in China. Meanwhile, important achievements have been obtained. Recently, it is proposed to upgrade the HIRFL with a multifuntional Cooling Storage Ring (CSR) forming a HIRFL-CSR accelerator system. This will greatly enhance the performances of the HIRFL for the first decade of next century, particularly for studies by using Radioactive Ion Beams(RIBs), and to meet the arising nuclear physics needs.

CSR, a new accelerator planned at the Heavy Ion Research Facility in Lanzhou (HIRFL), is a multipurpose Cooling Storage Ring system, and consists of a main ring (CSRm) and an experimental ring (CSRe). The existing HIRFL will be used as its injector system. The heavy ion beams with energy range of 10-100MeV/u from the HIRFL will be accumulated, cooled and accelerated to the higher energy range of 100-400MeV/u in the main ring, and then extracted fast to produce radioactive ion beams (RIBs) or highly charged state heavy ions. After that, the secondary beams (RIBs or highly charged ions) can be accepted by the experimental ring for many internal target experiments with cooling.

Two electron coolers located in the long straight sections of the CSRm and CSRe respectively, will be used for beam accumulation, compensating for the growth of beam emittance during internal target experiment and providing high quality beams for special experiments.

The CSR system intends to provide internal and external target beams for many physics experiments. Two internal target in the long straight sections of CSRm and CSRe will be used for nuclear physics and highly charged atomic physics. Two external target of CSRm will be used for cancer therapy studies and the researches on the properties of nuclear matter under extreme condition.

The main parameters of the CSR are listed in Table 1.The lattices of the CSRm and CSRe are quite similar to each other. They consist of four arc sections and two long straight sections. The doublet and triplet quadruple are the essential options.

Table 1. Main parameters of the CSR

CSRm	CSRe

____________________________________________________________________

Circumference(m)	161.20	128.96

Average radius(m)	25.66	20.56

Max.energy(MeV/u)	900(12C6+)	400(12C6+)

400(238U72+)	250(238U90+)

Bpmax/Bpmin(T-m)	10.64/1.25	6.44/0.64

Bmax/Bmin(T)	1.40/0.16	1.40/0.13

Ramping rate(T/s)	~10

______________________________________________________________________

The combination of multiturn injection and RF stacking will be used for the CSRm to accumulate heavy ions up to 109-7 in a short duration of about four seconds. During the accumulation, both electron and stochastic cooling will be used for cooling of beam phase space in order to increase the accumulation ratio and efficiency. Table 2. is the accumulation parameters of several typical ions.

Table 2. Parameters of accumulation in CSRm

12C6+	129Xe36+	238U37+

--------------------------------------------------------------------------------------

Energy(MeV/U)	50	20	10

Current(pps)	3.1x1012	8.7x1010	8.5x109

Stacking cycle(ms)	30	20	20

Injection period(s)	 4	 4	4

Particles of accum.	2.4x109	1.5x108	2.1x107

--------------------------------------------------------------------------------------

For the HIRFL-CSR plan, all optimising design and study are now in progress. The total budget of the CSR is about 300 million Chinese Yuan. The whole duration of it is intended to be seven years, two years for design and five years for construction. Zengquan Wei, IMP, CAS, Nanchang Road 363, Lanzhou 730000, P.R.China

News from St Petersburg, Russia: Pulsed dose rate of high-energy proton beam as DNA structural damages modifier.

Two regimes of GatchinaÕs synchrocyclotron 1000 MeV proton beam generated with frequency of macropulses 4,5 s-1 accompanied by pulsed dose rate 25 Gy s-1 (regime I) and with frequency 45 s-1 accompanied by pulsed dose rate 2,5 Gy s-1 (regime II) were compared using normal rats, RL-67 adenocarcinoma bearing mice (C57BLxCBA)F1 and DNA solution as a biological targets. Absorbed doses were monitored with ionising chamber and by yield of reaction C12-p-p,n (C11((+, E=0,511 MeV, ( 20 min.). The total absorbed doses and total elapsed time of exposes remained the same for both compared regimes.

The slight increasing of proton beam efficiency at regime I (RBEI:RBEII ( 1,1 ( 1,2) was established earlier in terms of lethality after total irradiation (6,9 and 12 Gy) of rats, the loss of their body weight, hematocrit, common protein level in plasma as well as raising free aminoacids concentration in it (D. Karlin, B. Konov, 1989).

The size or calculated volume of carcinoma metastases in lungs of locally irradiated mice (10, 30 and 60 Gy) were more diminished at regime I, then at II one (II : (I ( 2 or VI:VII ( 8, p=0,03).

The intrinsic viscosity [(] of DNA, the yield of 3'OH-ends and malon aldehyde (MA)-like ends after irradiation at doses 30-90 Gy were estimated by means of magnetic ratator viscometer, liquid-spectrometer ÒPicker-NuclearÓ, DEAE-chromatography, labeling 3'OH kit N67 (Amersham) and reaction DNA-MA-like end +NH-labelled AdR* as well as NH-GdR* or NH-CdR* (TdR* for lack of NH-groups was employed as a control). Statistical analysis of complete data showed the significant increasing of 3'OH-and MA-like ends/mkg DNA after irradiation could be demonstrated with regime II only. Partial restoration of [(] value for double strand DNA as well as for melted single strand DNA both registered in time after exposure at regime I. Therefore it is concluded from this study that capability of irradiated DNA in solution to form the cross-links between its damaged deoxyribose moiety of one strands and undamaged base of another strand has depended just on the value of pulsed dose rate parameter. The proposed appraisal of radiochemical output of lethal structural damages of DNA like the cross-links seems to be the fruitful approach for elucidation the nature of radiobiological effects of fractionated high-energy particles. A. Shoutko, D. Karlin, B. Konov, N. Chatinina, M. Vasilyeva, L. Ekimova, Central Research Institute of Roentgenology and Radiology, Leningradskaja str. 70/4. Pesochnij-2, St.-Petersburg, Russia, 186646.

A very compact Protontherapy facility based on an extensive use of High Temperature Superconductors (HTS).

The present worldwide protontherapy development is slowed down by the volume and the general costs of the existing proton generators and delivery dose equipments. The use of High Temperature Superconducting cables recently available on the market (HTS), working around 30K modify completely the design of protontherapy equipment. At this temperature, independant cryocoolers able to work in any position can be extensively used. Added to a significant simplification of the associated cryothechnics, it becomes possible to gain what was promised by the use of superconductivity concern ing the volume, the weight, the general dimensions, the consumpted power. The complete facility is composed of a very compact isochronous Cyclotron 210.Mev p of only 88 tons 3.2m in diameter equipped with an ECR ion source of infinite lifetime, followed by a superconducting voxel scanned isocentric gantry of 25 tons only, 5.2 meters in diameter. C. Bieth, A. Laisne, Pantechnik 14 rue Alfred Kastler, 14000 Caen, France; P. Mandrillon, Centre Antoine Lacassagne, 168 Avenue de la lanterne, 6000 Nice, France; Dr Ponvert, Institut Curie, 26 Rue d'Ulm 75248 Paris Cedex 05, France.

Guest Commentary by M. R. Raju

Dear Friends,

As a graduate student in nuclear physics in India during the late 1950Õs, I was fascinated by and enjoyed reading about neutron capture theory and Bragg peak therapy in the proceedings of the First United Nations Conference on Peaceful Uses of Atomic Energy. I had the good fortune to pursue a research career in this field with the help of Prof. Gordon Brownell at the Massachusetts Institute of Technology and Massachusetts General Hospital in Boston and John Lawrence (brother of Ernest O. Lawrence) at the Radiation Laboratory in Berkeley.

I never expected that I would have the privilege of being part of the continuing development of particle therapy for more than three decades and my last two publications were invited review papers. The first in Radiation Research, ÒParticle therapy: Historical Developments and Current StatusÓ 145, 391-407 (1996) and the second in Int. J. Rad. Biol., ÒProton Radiobiology, Radiosurgery and RadiotherapyÓ, 67, 237-259 (1995). I am happy that I can follow the progress in particle therapy through the Particles newsletter, which is an excellent informal journal. Over the years, I enjoyed my friendship with many of you, and I would like to take this opportunity to express my deep appreciation.

The purpose of this letter is to keep you all informed of my new venture and to seek your help and guidance, so that I can continue to interact with those of you who are interested in my new project. Nearly four years ago, I took voluntary retirement to spend most of my time developing a rural cancer center which will serve a mostly illiterate population in a rural area of Andhra Pradesh. The major emphasis of this cancer center will be on cancer awareness, prevention, early detection and treatment of the needy, who cannot afford to travel to big cities or pay the treatment costs. The appropriate technology needed, includes a Cobalt-60 treatment facility, a good pathology laboratory and a surgical theatre. I am already in the process of building the surgical theatre, butI seek your help in finding resources to acquire a Cobalt-60 treatment facility.

This rural cancer center will be a small step in overcoming the gulf between sophisticated medical research, where there is a lot of effort to develop new ways of improving treatment, and the needs of most people in the world, who can gain enormous benefits from relatively simple medical measures. The ÒMahatma Gandhi Memorial Medical TrustÓ, founded nearly 20 years ago by my wife and occupying nearly 15000 square feet of buildings located on a five acre site in a rural area is making my work much easier. In the future, I would like to invite radiation oncologists and medical physicists to visit the Trust periodically to serve the poor, and personally experience the living conditions of most people in the world. We have already built a guest house our friends to stay in, when they come to participate in the services of the Trust and its efforts.

I have been very fortunate to have the enthusiastic support of some of the leading scientists and radiation oncologists from India. With a grant from the Indian Atomic Energy Regulatory Board for Cancer surveillance and the treatment of cervical cancer using intracavity brachytherapy, we have been able to treat 16 patients and complete a door-to-door survey of about 25000 families in this rural area. We found that nearly 10% of the deaths were due to cancer and that this percentage is increasing. We also found that most cervical cancer patients are illiterate. Only about 10% of the population seek qualified medical help, the rest are in the hands of quacks. More than 80% of the cancer patients seek treatment at advanced stages of the disease and this number is not increasing yet.

Recently, I was asked to be the Member Secretary for a District Cancer Control Program, under the National Cancer Control Program which has support from the World Health Organization. The major emphasis of this program is on cancer awareness, prevention and early detection. We are developing pamphlets in local language on cancer prevention and early symptoms including self examination of the breast. Cervical, oral and breast cancers together comprise nearly 75% of the cancer problem in India. Cervical cancer among women and oral cancer among men are the two most common cancers in India; both these tumors are preventable and can be detected at even precancerous stages. However, the incidence of breast cancer seems to be increasing. With the help of local doctors in villages, we are conducting cancer screening camps to detect early stages of cancer and then we help these patients to get the necessary treatment in cancer centers.

Some medical problems in India are due to illiteracy, but it still surprised me to learn that nearly 75% of children drop out during elementary school. Another pet project of mine is to develop preschools in villages to benefit children of illiterate parents. In collaboration with local people, we have established such a school in one village and it is quite successful. My wife and I are now building a model preschool in another village on a site that I inherited, and then I hope to see such schools in as many villages as possible. I am hoping that these preschools will reduce the drop out rate during elementary schools, and the State Government seems to be interested in partially funding such schools.

If you are interested in visiting us, please write to me. At this point I have neither fax or electronic mail. Also, if you happen to be traveling in India, you are welcome to come and visit us. My wish is that some of the people working in the forefront of radiotherapy developments will also take some time to address some of the basic problems, such as those I have outlined here. M. R. Raju, Mahatma Memorial Medical Trust, Bhimavaram (AP) 534202, India. Telephone +91 8816 231 67.

Proposed NEW FACILITIES for PROTON & ION BEAM THERAPY

July 1998

| INSTITUTION | PLACE | TYPE | 1ST RX? | COMMENTS |
| --- | --- | --- | --- | --- |
| NPTC (Harvard) | MA USA | p | 1998 | at MGH; 235 MeV cyclotron; 2 gantries + 3 horiz. |
| Kashiwa | Japan | p | 1998 | 235MeV cyclotron;2gantries;1horiz; under construction |
| INFN-LNS, Catania | Italy | p | 1999 | 70 MeV; 1 room, fixed horiz. beam |
| Bratislavia | Slovakia | p, ion | 2000 | 75 MeV cyclotron; p; ions; +BNCT, isot prod. |
| CGMH, Northern Taiwan | Taiwan | p | 2000 | 250 MeV synchrotron, 3 gantries, I fixed beam |
| Hyogo | Japan | p, ion | 2001 | 2 gantries; 2 horiz; 1 vert; 1 45 deg;under construction |
| NAC, Faure | South Africa | p | 2001 | new treatment room with beam line 30o off vertical. |
| Tsukuba | Japan | p | 2001 | 270 MeV; 2 treat rooms with gantries; 1 research room |
| Wakasa Bay | Japan |
| 2001? | multipurpose accelerator; building completed mid 1998 |
| Shizuoka Cancer Center | Japan |
| 2002? | synchrotron 230? MeV; 2 gantries; 1 horiz; funded. |
| CNAO, Milan & Pavia | Italy | p, ion | 2002? | synchrotron; 1 gantry;2 fixed beam rooms;1 exp. room |
| AUSTRON | Austria | p, ion | ? | 2p gantry;1 ion gantry;1 fixed p;1 fixed ion;1 exp room |
| Beijing | China | p | ? | 250 MeV synchrotron. |
| Central Italy | Italy | p | ? | cyclotron; 1 gantry; 1 fixed |
| Clatterbridge | England | p | ? | upgrade using booster linear accelerator. |
| TOP project ISS Rome | Italy | p | ? | 200 MeV linac;1eye room; gantry?;under construction |
| 3 projects in Moscow | Russia | p | ? |
|  | HIRFL,Lanzhou | PR China | C ion | ? |
|  | JŸlich (KFA) | Germany | p | ? | exp. beam line; ? plans for therapy. |
| Krakow | Poland | p | ? | 60 MeV proton beam. |
| KVI Groningen | The Netherlands | p | ? | plan:- 200 MeV accel.; 2 rms; 1 gantry; 1 fix. |
| Moscow | Russia | p | ? | 320 MeV; compact, probably no gantry |
| Proton Development N.A. Inc. | IL USA | p | ? | 300 MeV protons; therapy & lithography |
| PROTOX | England | p | 2001? | existing RAL synchrotron; 250 MeV; 3 treat. gantry |

WORLD WIDE CHARGED PARTICLE PATIENT TOTALS

July 1998

| WHO | WHERE | WHAT | DATE | DATE | RECENT | DATE |
| --- | --- | --- | --- | --- | --- | --- |
|  | FIRST | LAST | PATIENT | OF |
|  | RX | RX | TOTAL | TOTAL |
| Berkeley 184 | CA. USA | p | 1954 | Ñ 1957 | 30 |  |
| Berkeley | CA. USA | He | 1957 | Ñ 1992 | 2054 | June-91 |
| Uppsala | Sweden | p | 1957 | Ñ 1976 | 73 |
|  | Harvard | MA. USA | p | 1961 |
| 7694 | Jul-98 |
| Dubna | Russia | p | 1967 | Ñ 1974 | 84 |
|  | Moscow | Russia | p | 1969 |
| 3039 | Dec-97 |
| Los Alamos | NM. USA | ¹- | 1974 | Ñ 1982 | 230 |  |
| St. Petersburg | Russia | p | 1975 |
| 1029 | Jun-98 |
| Berkeley | CA. USA | heavy ion | 1975 | Ñ 1992 | 433 | June-91 |
| Chiba | Japan | p | 1979 |
| 96 | Oct-96 |
| TRIUMF | Canada | ¹- | 1979 | Ñ 1994 | 367 | Dec-93 |
| PSI (SIN) | Switzerland | ¹- | 1980 | Ñ 1993 | 503 |
|  | PMRC, Tsukuba | Japan | p | 1983 |
| 576 | Mar-98 |
| PSI (72 MeV) | Switzerland | p | 1984 |
| 2487 | Dec-97 |
| Dubna | Russia | p | 1987 |
| 40 | May-98 |
| Uppsala | Sweden | p | 1989 |
| 147 | Feb-98 |
| Clatterbridge | England | p | 1989 |
| 817 | May-98 |
| Loma Linda | CA. USA | p | 1990 |
| 3433 | Apr-98 |
| Louvain-la-Neuve | Belgium | p | 1991 |
| 21 | Nov-93 |
| Nice | France | p | 1991 |
| 1010 | Jan-98 |
| Orsay | France | p | 1991 |
| 956 | May-97 |
| N.A.C. | South Africa | p | 1993 |
| 243 | Dec-97 |
| IUCF | IN USA | p | 1993 |
| 1 | Dec-97 |
| UCSF - CNL | CA USA | p | 1994 |
| 162 | May-98 |
| HIMAC, Chiba | Japan | heavy ion | 1994 |
| 389 | Feb-98 |
| TRIUMF | Canada | p | 1995 |
| 37 | Jan-98 |
| PSI (200 MeV) | Switzerland | p | 1996 |
| 9 | Dec-97 |
| G.S.I Darmstadt | Germany | heavy ion | 1997 |
| 2 | Dec-97 |
| Berlin | Germany | p | 1998 |
| 3 | Jul-98 |
|  | 1100 | pions |
|  | 2878 | ions |
|  | 21987 | protons |
| TOTAL | 25965 | all particles |

GGGGGGGGGGGGGGGGGGGGGGGG	GG

G                                                                          	G

G                            See Page 19.                  	G

G                                     for                             	G

G     The Proposed New Facilities Table 	 G

G                                                                                	G

GGGGGGGGGGGGGGGGGGGGGGGG	GG